# Embedded Code Guide – rev A

December 22, 2015

Megan Hodgson
Infinite Biomedical Technologies
megan@i-biomed.com

# Table of Contents

Table of Contents	2

Purpose	3

Included Files	4

Source Files (src)	4

Include files	6

Libraries	7

Software Flow	8

General Flow	8

Specific Flow	9

Table 1: Keyboard code key options	12

Notes:	13

# Purpose

The purpose of this document is to guide the user through use of the:

- Keyboard code to control the Touch Bionics robolimb CAN hand and Motion Control direct drive wrist motor

- Direct Control code for use with IBT electrodes

Use with code version--
IBT-CORE, robolimb_keyboard branch
Commit version: 23 Dec 2015, 15:09

For use with dev board rev B 1.0

# Included Files

K = Used in keyboard mode, D = Used in direct mode

### Source Files (src)

config.cpp (K/D)

- Purpose: controls saving and loading of configuration information (run mode, number of electrodes, etc.) to/from the flash memory internal to the ARM

- Status:

  - Tested and working (only some information currently used)

  - Information saved in configuration can be modified according to needs

- Important functions:

  - Config_get_configuration: load last configuration

  - Config_set_configuration: update memory with new configuration

data.cpp (D)

- Purpose: saves a circular data buffer to compute the electrode envelope value and computes that output value

- Status:

  - Tested and working

- Important functions:

  - Data_add_data: called within read_sensor_data interrupt to add new data to the circular data collection buffer

  - Data_get_envelop: called within RUN mode, when in Direct control mode, to compute the envelope output value

electrode.cpp (D)

- Purpose: includes methods for initializing, communicating with and gathering data from the ADS1291 AFE on the IBT electrode

- Status:

  - Tested and working

- Important functions:

  - Electrode_init: controls the initialization of the electrode, including gains (see Notes)

  - Electrode_read_data: collects a data point from the electrode within the read_sensor_data interrupt

filter.cpp (D)

- Purpose: includes functions to filter new electrode data to remove ambient noise

- Status:

  - Tested and working

- Important functions:

  - HighPass_filter_IIR: first level high pass filter (cut-off frequency options 10Hz, 15Hz, or 20Hz)

  - Integral_High_Pass_Filter: second level high pass filter

  - Notch_filter_IIR: narrow notch filter to remove powerline noise

h_bridge.cpp (K)

- Purpose: includes function to interact with the motor driver board

- Status:

  - Tested and working

- Important functions:

  - HB_reverse: enables PWM in one direction

    - Second parameter should be between 0 and 1 (1 = 100% duty cycle)

  - HB_forward: enables PWM in other direction

    - Second parameter should be between 0 and 1 (1 = 100% duty cycle)

  - HB_brake: stops the wrist movement

log.cpp (K/D)

- Purpose: used to log events in the code, especially configuration changes and errors. Information saved to internal flash memory

- Status:

  - Tested and works

- Important functions:

  - Log_entry: parameters are “event type”, “number of data bytes” and “data”; the information is logged into the next  available location in memory

main.cpp (K/D)

- Purpose: controls the overall functionality of the program

- Status:

  - Tested and working

- Important functions:

  - See Software Flow Section of document

MCP2515.cpp (K)

- Purpose: to communicate with the external  CAN controller

- Status:

  - Sending functions tested and working

  - Receiving functions are still in testing – interrupts and flags need to be enabled

- Important functions

robocan.cpp (K)

- Purpose: includes functions to control the CAN hand

- Status:

  - Sending functions tested and working

  - Receiving function are still in testing (labeled as such in the code)

- Important functions:

  - RC_send_grip_command: updates the grip on the hand

  - RC_send_digit_command: updates the direction/speed of specified digit

### Include files

ads1291.h (D)

- Includes settings and commands for the various registers of the AFE chip on the electrode (ADS1291)

config.h (K/D)

- Includes initializations for configuration structure and associated functions in config.cpp

data.h (D)

- Includes initializations for functions in data.cpp

electrode.h (D)

- Includes initializations for functions in electrode.cpp

errors.h (K/D)

- Lists the error codes for problems that may arise while running the program (when rc != LDA_OK)

filter.h (D)

- Includes initializations for public functions in filter.cpp

h_bridge.h (K)

- Includes initializations for functions in h_bridge.cpp

input_data.h (--)

- Included so that config.cpp will compile (not used)

log.h (K/D)

- Includes initialization of structure for log entries and functions in log.cpp

MCP2515.h (K)

- Defines the registers for the MCP2515 and associated control functions

robocan.h (K)

- Defines the commands to control the robolimb CAN hand as described in the manual

types.h (K/D)

- Defines new type definitions used in throughout the program

### Libraries

- mbed-rtos: schedules processing

- mbed-src: includes common libraries such as Serial, InterruptIn, etc.

- storage_on_flash: communicates with the on-board flash

# Software Flow

### General Flow

CONFIG

Setup the type of control system to run the rest of the software

INIT

Initializes the peripherals based on the configuration

BIT

Checks that peripherals initialized correctly

IDLE

Waits to let the buffers fill/code settle

RUN

Puts the code into normal operation

ERROR

Enters this mode with error in another mode; logs and returns to BIT

### Specific Flow

- Program start

- CONFIG runs:

  - If you would like to change the configuration, press ‘y’ or ‘Y’

    - Choose control system:

      - For DIRECT CONTROL, press ‘1’ and ‘ENTER’

      - For KEYBOARD CONTROL, press ‘2’ and ‘ENTER’

      - The serial output will list the mode that was configured at the end of CONFIG mode

  - If you would like to keep the last configuration, press ‘n’ or ‘N’

- INIT runs:

  - Turn on power to hardware peripherals

  - If in DIRECT mode:

    - Resets process flag to 0

    - Initialize the filter properties

    - Initialize the interrupt to read data from electrodes every 1msec

    - Initialize 8 electrodes (may take up to 2 minutes)

    - Initialize data collection buffers for each electrode

  - If in KEYBOARD mode:

    - Sets process flag to 1

    - Initialize the interrupt to receive new data from the keyboard

    - Set output selector switch and initialize CAN controller

    - Initialize H-bridge 2

  - Initialize Logging functions

    - To save logs, set answered to ‘n’ at the point in the code after the printf “Erase Logs”

    - Serial output should say “Hardware Initialized!” at the end of INIT

- BIT runs:

  - If in DIRECT mode:

    - Run electrode BIT – should read ID of electrode

  - If in KEYBOARD mode:

    - Run H-bridge BIT – wrist should make short turns in both directions

    - Run CAN hand BIT – hand should full open and fully close

- IDLE runs:

  - Serial output should output “Waiting for a few seconds.” And then be ready for use

  - Note: if there is an “ERROR!! 0” after IDLE for the DIRECT mode, power cycle the board and run the software again.

- RUN runs:

  - If in DIRECT mode:

    - If process = 0, code will not run.

    - Every 1msec, timer interrupt fires, and for each electrode:

      - Collect Electrode data from ADS1291

      - Filter electrode data (notch and high pass)

      - Input new data to buffers

      - Set process = 1 to enter DIRECT mode in RUN mode in main

    - Now, process = 1

    - Apply envelope algorithm to each electrode separately

    - Every 10msec, all electrode values are output to the serial port as integers

NOTE: If using the C# GUI to view the electrode signals, it can only display 1 electrode at a time. There is a different sending structure so that the GUI can recognize the data: #, High data byte, Middle data byte, Low data byte, $. To send only 1 electrode, connect to Tera Term and hit the # of the electrode you would like to output. See Notes section to verify which key corresponds to which electrode (depends on which armband you are using). If you wish to return to output all signals, hit ‘a’.

Interrupt at 1msec:

- Collect data point from each electrode

- Filter data and add to data buffer

- Set process = 1

Is process = 1?

Enter RUN mode

Apply envelope to all electrode signals.

Output all 8 electrode signals

OR

Output 1 specific electrode signal in GUI compatible format

Process = 0

  - If in KEYBOARD mode:

    - If new_cmd = 0, code will not run.

    - If new key was pressed, interrupt fires:

      - Use getc to get the new character (key_cmd)

      - Set new_cmd = 1

    - Now, new_cmd = 1

    - Apply key_cmd value to a switch statement

    - When the key command is a Hand Grip commands:

      - Stop hand and wrist movement

      - Check if grip command has already been sent

        - Note: if a grip command is sent previously, sending it again puts the hand in normal mode

      - If not, send the new grip command

    - When the key command is open/close hand:

      - Stop wrist movement

      - Send command to open or close all digits

      - Only fingers that are active in that grip will move

      - Set finger directions (0 = open, 1 = close)

    - When the key command is to control individual fingers:

      - Stop wrist movement

      - Check the last direction that the chosen finger moved

      - Switch the finger direction flag and move the finger the opposite direction

    - When the key command is to control the wrist:

      - Stop hand movement

      - Move wrist indefinitely

    - When a Stop command is called:

      - There are two commands:

        - Stop all movement

        - Stop wrist movement

- In ERROR mode:

  - Errors will print out to Serial port

  - Check errors.h for meaning of Error values

### Table 1: Keyboard code key options

| Key Pressed | Action |
| --- | --- |
| 0 | Normal grip |
| 1 | Standard Precision Pinch Closed |
| 2 | Standard Tripod Closed |
| 3 | Thumb Park, continuous |
| 5 | Lateral grip |
| 6 | Index Point (trigger) |
| 7 | Standard Precision Open |
| 9 | Thumb Precision Closed |
| a | Thumb Precision Open |
| b | Thumb Tripod Open |
| d | Standard Tripod Open |
| e | Thumb Tripod Open |
| f | Donning/doffing glove |
| o | Open all fingers (depends on active fingers in that grip) |
| c | Close all fingers (depends on active fingers in that grip) |
| t | Move thumb individually |
| i | Move index finger individually |
| m | Move middle finger individually |
| r | Move ring finger individually |
| p | Move pinky finger individually |
| s | Stop all movement |
| x | Stop movement of wrist |
| y | Move wrist one direction |
| z | Move wrist other direction (depends on cable orientation) |

Figure: the output from the serial port when in Keyboard Mode

# Notes:

- For keyboard code, the hand selection chip must be on the write output to communicate with the CAN hand: handout_sel.write(0)

- The electrodes have 7 discrete gain values for the initial ADC step: 1, 2, 3, 4, 6, 8, and 12. A gain of 6 is the default. In order to modify this value, you must change the gain indicator in lines 171 and 178. The other defined values for gain options are listed in ads1291.h, lines 65 to 71

- To select 50Hz or 60Hz powerline noise removal, choose the correct filter coefficient options in Filter_init

- The CAN controller must be in CONFIGURATION mode in order to change settings. Use setMode function in MCP2515.cpp

- Must include types.h in any file that you want to use pcserial output

- To intercept raw electrode signals, use the output of Filter_process_sensor_input

- The two armbands have different CS orders to make initialization more efficient for each one

  - Armband 1

| Key | CS pin | Electrode # (1 is the electrode where the SPI/power lines start, 8 is where they end) |
| --- | --- | --- |
| 1 | 8 (PA_12) | 5 |
| 2 | 3 (PC_12) | 2 |
| 3 | 6 (PC_5) | 7 |
| 4 | 2 (PC_11) | 3 |
| 5 | 1 (PA_0) | 1 |
| 6 | 4 (PA_4) | 4 |
| 7 | 5 (PA_8) | 8 |
| 8 | 7 (PC_13) | 6 |

  - Armband 2

| Key | CS pin | Electrode # (1 is the electrode where the SPI/power lines start, 8 is where they end) |
| --- | --- | --- |
| 1 | 1 (PA_0) | 1 |
| 2 | 2 (PC_11) | 2 |
| 3 | 3 (PC_12) | 3 |
| 4 | 4 (PA_4) | 4 |
| 5 | 5 (PA_8) | 5 |
| 6 | 6 (PC_5) | 6 |
| 7 | 7 (PC_13) | 7 |
| 8 | 8 (PA_12) | 8 |